## Read Me
#REF!
The Tilt test workbook provides a simplified procedure to fine tune the vertical sensitivity axis of an FBV seismometer.
This test emulates the tilt-table testing used on professional instruments, but uses shims of varying thickness
placed under the instrument legs to generate known amounts of tilt. It is performed on an operational instrument, and does
not require a sealed pressure case. A simple draft shield should be adequate.
Measurements are made of the Centering Force Output voltage, starting with the instrument leveled, then with varying
shim thicknesses under the single leg or under the two legs at the coil end. It will require about 5 minutes for the voltage
to settle each time. To speed up the process, R5 may be temporarily paralleled with a resistor of roughly 1/10 its value.
The measured data are entered in the yellow cells of the spread sheet, which corresponds to the number of data points
available. Results will be in cells C29-C32.
The worksheets are designed to analyze measurements using one, two or three shim thicknesses, using 3, 5, or 7 data
points. A single shim, providing three data points, should be sufficient to align the vertical axis accurately enough for
normal use. Using more shims and more data points allows for more precise adjustment.
Suggested shim dimensions for the Inyo are 4.5" x 1" and for the Yuma 4" x 1", though if the working surface is not
perfectlly flat, using pairs of 1" x 1" shims may be preferable. Thicknesses of 1/4" 1/8" and 1/16" should be useful,
though anything in that general range will work. The maximum total shim thickness will be limited by having to maintain the
Centering Force voltage within ±14V. (±10V if measuring with 10V A/D)
The applicable instrument parameters should be entered in cells C3-C10, where XL is the X-axis distance between the
legs. All except for XL are not used at present but might be used in the future to derive other instrument parameters.
If desired, a more precise value for local 'g' could be entered in cell I3.
The chart which corresponds to the worksheet used will display a fitted parabolic approximation to the actual cosine
curve, which results from instrument tilt. The worksheet locates the maximum or minimum point on that curve, which will
be the angle of the vertical sensitivity axis deviation. To correct that angle, a trim mass may be raised or lowered until the
angle is near enough to zero degrees. The sample curves assume the coil polarity is '+ = Up'. If it was installed '- = Up'
the plot will appear inverted from the sample and also the response to horizontal mass adjustments will be reversed.
In practice, a degree or two deviation from vertical is not a serious problem, though good commercial instruments ususually
specify vertical alignment accuracy of ~0.3 degrees (0.5% cross-axis sensitivity). However, if you are planning to estimate
instrument noise by observing the coherence in the signals from two or three instruments, considerable care in aligning the
sensitivity axes of the instruments will be required.
This procedure also provides information which will be useful for adjusting the horizontal trim weight. A perfectly
aligned instrument will have the peak of the curve occuring at 0 degrees tilt and 0 Volts of Centering Force signal, though
it should be noted that the CF voltage will change somewhat in response to temperature changes.

## 7_Points
Tilt_Test_357.xls |  | #REF!
|  |  |  | Inyo | Yuma | Napa
| GN = | 9.6 | N/A | 15.5 | 10 |  |  | 9.7998 | N/kg | = local 'g'
| Ri = | 15000 | Ω | 12000 | 15000 |  |  | From | http://www.ptb.de/cartoweb3/SISproject.php
| Rc = | 43 | Ω | 50 | 44
| R22 = | 100000 | Ω | 100000 | 100000
| R7 = | 100000 | Ω | 100000 | 100000
| R25 = | 100 | Ω | 1000 | 100
| RLOAD = | 100000 | Ω | 100000 | 100000 | 100000
| XL = | 4.875 | in | 7.5 | 4.875
Spacers under coil end - positive Zs
Spacer under pivot end - negative Zs
| Measured |  | θ's to fit |  |  | Fitted Parabola
| Spacer Zs - in | Cent. Force VCF - Volts | Tilt θ Deg. | θ2 |  | θ to plot | VCF to plot |  | Fitted Parabola VCF = a2θ2+a1θ+a0 | a2 | a1 | a0
| -0.375 | 4.258 | -4.4117 | 19.4633 |  | -5.2941 | 6.0674 |  | Values | 0.22594 | 0.11513 | 0.34453
| -0.25 | 1.9334 | -2.9395 | 8.6409 |  | -4.2353 | 3.9096 |  | Std. Errors | 0.0020599 | 0.0052487 | 0.023568
| -0.125 | 0.6565 | -1.4693 | 2.1588 |  | -3.1764 | 2.2585 |  | r2 and Std Err in Vi | 0.99968 | 0.040834 | #N/A
| 0 | 0.3078 | 0 | 0 |  | -2.1176 | 1.1139 |  | F, Deg of Freedom | 6255.5584 | 4 | #N/A
| 0.125 | 1.042 | 1.4693 | 2.1588 |  | -1.0588 | 0.47592 |  | Sum Sqs. | 20.8613 | 0.0066697 | #N/A
| 0.25 | 2.672 | 2.9395 | 8.6409 |  | 0 | 0.34453
| 0.375 | 5.217 | 4.4117 | 19.4633 |  | 1.0588 | 0.71972 |  | Named Ranges | Cells
[6 empty] | 2.1176 | 1.6015 |  | VCFForFit | C16:C22
[6 empty] | 3.1764 | 2.9899
[6 empty] | 4.2353 | 4.8848 |  | θForFit | E16:F22
[6 empty] | 5.2941 | 7.2864
As viewed from the side, with the coil on the right, pivot on the left.
| max/min VCF | 0.32986 | V
| θ0 @ max VCF | -0.25479 | deg | Sensitivity angle CCW from vertical (+ if Vert. trim wt. too high)
| H sensitivity = | 0.4447% | of Vertical
| F(θ0) = | -0.021481 | gm | Centering Force error. (+ if Horiz. Trim wt. too far from pivot) Reduce M0 by this amount by reducing or moving Horizontal trim wt towards pivot, or away if negative.
-1.59
-1.0179
-0.5727
-0.25458
-0.06365
0
-0.06365
-0.25458
-0.5727
-1.0179
-1.59