## Sheet1
Loop testing - Yuma2 Rev 4.0
5/9/2015 Brian Kuschak <bkuschak@yahoo.com>
See http://bnordgren.org/seismo/FBV/Loop_Test_Board-2.pdf
Loop test board
Sig gen
Oscope
Probe calibration: Vout channel is reading 3.2% low. Use Vout (corrected) column
Loop phase measured two ways. Directly by measuring phase shift between Vout and Vin with scope, and calculated using voltages.
Osc Freq (Hz) | Vosc (pp) | Vin (pp) | Vout (pp) | Phase shift (s) | Vout (corrected) | Loop gain (dB) | Loop phase (deg) | Loop phase (deg) | Comments
1 | 6.4 | 0.066 | 6.24 | 0.224 | 6.4459 | 39.7948 | 80.64 | 45.7014 | vin accuracy low
2 | 5 | 0.17 | 4.92 | 0.11 | 5.0824 | 29.5123 | 79.2 | 60.1805
5 | 2.04 | 0.18 | 2 | 0.044 | 2.066 | 21.1972 | 79.2 | 79.2166
10 | 2.04 | 0.36 | 2.02 | 0.022 | 2.0867 | 15.263 | 79.2 | 77.6183
20 | 2.04 | 0.74 | 2.08 | 0.01 | 2.1486 | 9.2586 | 72 | 71.6211
30 | 2.04 | 1.1 | 2.2 | 0.0064 | 2.2726 | 6.3026 | 69.12 | 63.7272
40 | 2.04 | 1.6 | 2.3 | 0.0044 | 2.3759 | 3.4342 | 63.36 | 57.8721
50 | 2.04 | 2.12 | 2.46 | 0.003 | 2.5412 | 1.574 | 54 | 50.9335
57.57 | 2.04 | 2.6 | 2.6 | 0.0023 | 2.6858 | 0.28201 | 47.668 | 45.3681
60 | 2.04 | 2.76 | 2.62 | 0.0021 | 2.7065 | -0.17015 | 45.36 | 43.8107
70 | 2.04 | 3.44 | 2.76 | 0.0016 | 2.8511 | -1.631 | 40.32 | 36.3394
80 | 2.04 | 4.04 | 2.74 | 0.00104 | 2.8304 | -3.0906 | 29.952 | 28.1148
90 | 2.04 | 4.44 | 2.62 | 0.00076 | 2.7065 | -4.2996 | 24.624 | 17.8463
100 | 2.04 | 3.96 | 1.84 | 0.00012 | 1.9007 | -6.3755 | 4.32 | #NUM!
Results:
Crossover freq = 57.6 Hz
Phase margin = 47.7 deg